## Slide 1: Catastrophic Events
Michael F. Hancock, CCP

## Slide 2: Catastrophic Events
1/300 cases has an event
1/1000 patient dies due to an event
Types of Catastrophic Events
  Oxygenator failure
  Water to blood leak
  Air embolism
  Circuit Fracture
  Blood transfusion reaction
  Protamine reactions

## Slide 3: Oxygenator Failure
Oxygenators become less efficient at removing CO2 and adding O2
Due to manufacturers defect or long term use
  Usually a gradual failure over a few hours
  Can be acute due to thrombosis or fracture
  Oxygenators primed for an extended length of time can lose their efficiency
Signs-
  Decreased pO2
  Decreased venous saturation
  Elevated pCO2 (~ 80 – 100)
  Dark arterial blood

## Slide 4: Oxygenator Failure
Action- notify surgeon and anesthesia of possibility
  Check circuit for issues
    Trace gas pathway from source to oxygenator
  Increase oxygen delivery
    Increase gas flow
    Increase pump flow
    Increase BP
    Cool the patient more
  Take ABG to verify (after making changes)
If repeat ABG suggests failure:
  Quickly evaluate if you can finish the case or come off bypass
    Did you just go on CPB or are you close to being done?
      You never want to change out the oxygenator while the patient is on bypass especially if arrested

## Slide 5: Changing Out Oxygenator
2 Options:
  Traditional Oxygenator Changeout
  Entire Circuit Change (faster)
    Cut arterial and venous lines and connect to new pump
Oxygenator Changeout:
  Get extra oxygenator ready
  Come off bypass
  Change out oxygenator
  De-air new oxygenator
  Go back on bypass quickly

## Slide 6: Water to Blood Leak
Open pathway between the heat exchanger and the oxygenator, allowing the two fluids to mix
Due to fracture in the circuit
  Hook up water lines before priming to detect a water to blood leak
    You would see fluid in your resevoir before dropping prime
Signs-
  Sudden rise in blood volume
  Decrease in pH- (acidic)
  Decrease in HCT/Hbg= hemodilution
  Blood in urine (hemolysis)

## Slide 7: Water to Blood Leak
Action- notify surgeon and anesthesia
  Change out oxygenator
  Replace blood volume with donor blood and crystalloid
  Be prepared to correct acidosis

## Slide 8: Air Embolism
3 Types:
  Surgical Air
  CPB Air
  Anesthetic Air

## Slide 9: Surgical Air
Air emboli in the right side of the heart is well tolerated
  Expelled by the Lungs
Air emboli in the left side of the heart is often fatal
  Open left atrium, left ventricle, pulmonary vein, aorta
  Air is most commonly seen in the Right Coronary Artery during surgery
    Seen by Isolated Acute ST Segment Elevation in the inferior ECG leads

## Slide 10: CPB Air- (Pump Air)
Air via the arterial line due to emptying of CPB reservoir
Vortexing of blood at high pump flows through vertical ports of the venous reservoir cause air generation into the systemic arterial line
  Newer angled or horizontal outlets are less prone to vortexing
Arterial roller pump head tubing may rupture causing the entraining of air
  With a roller pump, if arterial pump is off and cardioplegia pump is on, air can be drawn into the plegia circuit
VAVD can deprime the arterial line filter or oxygenator if vacuum is applied before establishing CPB

## Slide 11: Treatment of Air Embolism-
With open chest, surgeon can directly aspirate air from the heart chambers or vessels
Air down the RCA is seen by ST Elevation on ECG
  Inducing HTN will help blow the air through the vessel
  Vasopressors can be given, surgeon can also pinch the aorta distal to the RCA to acutely raise BP to the RCA
Packing the head in ice and ventilating on 100% O2 can help
Head down position will help prevent bubbles from entering cerebral vasculature

## Slide 12: Massive Air Embolism-
Protective Measures against air embolism:
level sensors, arterial filters, bubble sensors protect against air emboli
CO2 flush, use albumin, keep adequate volume, communication
Actions-
immediately turn pump off and clamp your lines
notify surgeon
place patient in trendelenberg
pack head in ice
give steroids
carotid massage by anesthesiologist
drop fluid

## Slide 13: Massive Air Embolism
recirc fluid through recirc line and purge line to get rid of air
re-prime all necessary lines
if arterial filter is full of air, may have to clamp it out and just use bypass line to quickly re-initiate CPB
go back on bypass
if large quantities of air appear to have entered cerebral vessels, retrograde cerebral perfusion may be used with great results
hyperbaric chamber is pumped air to patient
  ~6 ATA (atmospheres of absolute pressure) is required to get rid of air

## Slide 14: Blood Transfusion Reactions-
Recipient forms antibodies against donor RBCs
Reactions can be immediate or delayed
Double check patient IDs on blood
Actions taken-
  stop transfusion
  give steroids
  give heparin
  give diuretics
  give a complete transfusion with the proper blood type
  fill out incident report for the hospital- (doesn’t go in patient chart)

## Slide 15: Protamine Reactions
Derived from salmon sperm
People at Risk-
  Causes reactions in people with shellfish allergy
  Risks when patient was already exposed to protamine
  Men with vasectomies
  Patients on NPH insulin

## Slide 16: Protamine Reactions
Types of Protamine Reactions-
  Type I- hypotension, due to fast administration of protamine
  Type II- hypotension, decrease in SVR, bronchospasm, edema, pulmonary vasoconstriction
  Type III- DEATH
    Pulmonary vasoconstriction, RV distention, PA HTN, decreased LV filling, systemic hypotension
Action taken-
  give vasoactive drugs
  give inotropic drugs
  give pulmonary vasodilators
  give heparin, go back on CPB